## Requirements
BD1818 Design Details - Double Beam Pecision Weighing Lysimeter
by Brad Lyles
1.8 | bulk density (gm/cm^3) |  |  |  | load cell calcs
0.3 | soil porosity (%) |  |  |  | 25 | max load (lbs) or FSR (Full Scale range)
0.35 | field capacity (%) |  |  |  | 11.3398 | max load (kg)
[5 empty] | 3.001 | mV/V FSR
[5 empty] | 2.5 | V excitation
[5 empty] | 7.5025 | mV FSR | without mechanical advantage
45.72 | diameter (cm) |  |  |  | 0.66161 | mV/kg | load cell calcs
18 | diameter (in) |  |  |  | 1511.4725 | g/mV | 1000 | max load (lbs)
22.86 | radius (cm) |  |  |  | 0.33 | uV analog measurement resolution | 0.02 | accuracy (+/- percent of FSR)
45.72 | h (cm) |  |  |  | 1.5115 | g/uV | 0.2 | accuracy (+/- pounds)
18 | h (in) |  |  |  | 0.49879 | resolution (g)
1641.7322 | area (cm^2)
75059.9977 | volume (cm^3) | 2.61 ft^3 | 4.35 | 60# sacks of sand | 11.3398 | max load (kg) | 453.5929 | max load (kg)
135107.9958 | soil mass (gm) |  |  |  | 0.003 | accuracy (+/- FSR) | 0.02 | accuracy (+/- FSR)
135.108 | soil mass (kg) | 297.8618 |  |  | 0.00034019 | accuracy (+/- kg) | 0.090719 | accuracy (+/- kg)
10 | mechanical mass (kg) |  |  |  | 0.34019 | accuracy (+/- gm) | 90.7186 | accuracy (+/- gm) | 1.13 | lysimeter resolution (g)
7.8813 | water mass at field capacity (kg) | 17.3753 | (lbs) |  | 6.8039 | theoretical accuaracy w/ 20x mechanical advantage (+/- gm) |  |  | 0.006883 | lysimeter resolution (mm)
152.9893 | total mass (kg) | 337.2833 |  |  | 0.041443 | theoretical accuracy flux (mm)
337.2833 | total mass (lbs)
| Flux Requirements |  |  |  | 1 | water bulk density (gm/cm^3) | 1 | water bulk density (gm/cm^3)
1 | water bulk density (gm/cm^3) |  |  |  | 0.34019 | accuracy as water(cc) | 90.7186 | accuracy as water(cc)
25.4 | flux requirement (mm) | 0.01 |  |  | 0.00020722 | flux(cm) | 0.055258 | flux(cm)
2.54 | flux requirement (cm) |  |  |  | 0.0020722 | flux(mm) | 0.55258 | flux(mm)
1641.7322 | area (cm^2)
4169.9999 | vol of water (cm^3)
4169.9999 | mass of water (gm) |  |  |  | 1 | (gm) resolution
[5 empty] | 0.00060911 | resolution flux (cm)
[5 empty] | 0.0060911 | resolution flux (mm)
Lyimeter mass mechanical advantage
152.9893 | total mass (kg) | 7.8813
8 | xxx times mechanical advantage | 20 |  |  | 3.9903 | +/- lysimeter resolution (g)
19.1237 | Max on load cell (kg)
42.1604 | Max on load cell (lbs)
24.7851 | Counter Wt needed for dry soil (lbs)
93.1877 | SWE (cm) | 36.6881

## Design as torque
Lysimeter design
| 1 | scale factor
Beam 1 | Transfers load from tank to connecting rod at Beam 2 | [8 empty] | Scaled
F1 | 152.9893 | Load (kg) |  |  |  |  | F1 | 1500.825 | Force (N) | 1500.825
F2 | 31.5819 | Wt at link (kg) |  | Orig | Final | Scaled | F2 | -300.5457 | Force (N) | -300.5457
d1 | 1 | Length to knife bearing (in) |  | 1 | 1 | 1 | d1 | 0.0254 | Length (m) | 0.0254
d2 | 5 | Length to link (in) |  | 7 | 5 | 5 | d2 | 0.127 | Length (m) | 0.127
M1 | -152.9893 | moment (kg-in) |  |  |  |  | t1 | 38.121 | Torque (Nm) | 38.121
M2 | 157.9093 | moment (kg-in) |  |  |  |  | t2 | 38.1693 | Torque (Nm) | 38.1693
M3a | -4.92 | moment from tubing mass | 4.9182 |  |  |  | tB1a | -0.0020145 | Torque (Nm)
[7 empty] | tB1b | 0.050364 | Torque (Nm)
Beam 2 | Transfers Wt at connecting rod to load cell
[8 empty] | 10.7251 | Load Cell Load (kg)
F3 | 31.5819 | wt at link from F2 (kg) |  |  |  |  | F2 | -300.5457 | Force (N) | -300.5457
F4 | 10.6637 | wt final on load cell (kg) | 23.5093 | (lbs) |  |  | F4 | 105.2135 | Force (N) | 105.2135
d3 | 3.5 | Length to knife bearing (in) |  | 1 | 3 | 3.5 | d3 | 0.0889 | Length (m) | 0.0889
d4 | 10 | Length to load cell (in) |  | 12 | 12 | 10 | d4 | 0.254 | Length (m) | 0.254
M3 | 110.5365 | moment (kg-in) |  |  |  |  | t3 | -26.7185 | Torque (Nm) | -26.7185
M4 | 106.6365 | moment (kg-in) |  |  |  |  | t4 | -26.7242 | Torque (Nm) | -26.7242
M4a | -3.9 | moment from tubing mass | 3.9 |  |  |  | tB2a | 0.00040291 | Torque (Nm)
[7 empty] | tB2b | -0.0061282 | Torque (Nm)
X factor | 14.3468 | mechanical advantage | [5 empty] | 14.2646 |  | 14.2646
Beam 1 tubing torque detail
|  | 2" square tubing mass | 4.6048 | lb/ft
|  | 2" square tubing mass | 0.63661 | kg/m
B1a_d | 0.5 | tubing length rt of fulcrum | 0.0127 | m
B1b_d | 2.5 | tubing length lt of fulcrum | 0.0635 | m
B1a_m |  | Mass rt of fulcrum | 0.01617 | kg
B1b_m |  | Mass lt of fulcrum | 0.080849 | kg
B1a_F |  | Force right of fulcrum | 0.15863 | N
B1b_F |  | Force left of fulcrum | 0.79313 | N
B1a_t |  | Torque right of fulcrum | -0.0020145 | Nm | [6 empty] | connecting rod matrix
B1b_t |  | Torque left of fulcrum | 0.050364 | Nm | [7 empty] | d2
[11 empty] | d3 | 5 | 6 | 7
[11 empty] | 2.5 | 20.3837 | 24.7173 | 29.1429
Beam 2 tubing torque detail | [10 empty] | 3.5 | 14.3468 | 17.343 | 20.3837
|  | 1" tubing mass | 0.92096 | lb/ft
|  | 2" square tubing mass | 0.12732 | kg/m
B1a_d | 0.5 | tubing length rt of fulcrum | 0.0127 | m
B1b_d | -1.95 | tubing length lt of fulcrum | -0.04953 | m
B1a_m |  | Mass rt of fulcrum | 0.003234 | kg
B1b_m |  | Mass lt of fulcrum | -0.012612 | kg
B1a_F |  | Force right of fulcrum | 0.031725 | N
B1b_F |  | Force left of fulcrum | -0.12373 | N
B1a_t |  | Torque right of fulcrum | 0.00040291 | Nm
B1b_t |  | Torque left of fulcrum | -0.0061282 | Nm
F1
d1
d2
F2
F2*d2 = F1*d1 - tB1a + tB1b
Level-one Beam
tB1b
tB1a
F2
tB2a
tB2b
d3
d4
F4*d4 = F2*d3 + tB2a - tB2b
F4
Level-two Beam

## Drainage Calcs
Assuming soil is at field capacitiy and all in coming precip yields 100% drainage.
1 | rain fall event (in)
2.54 | precip (cm)
1641.7322 | lysimeter area (cm^2)
4169.9999 | volume (cm^3)
4.17 | fluid volume (L) |  | sch 40
|  | pipe size | avg ID
4 | nominal pipe size | 2 | 2.047
3.998 | well dia (in) | 3 | 3.042
24 | well depth below inlet (in) | 4 | 3.998
10.1549 | well dia(cm) | 5 | 5.016
80.9921 | well area (cm^2) | 6 | 6.031
4937.2811 | well volume (cm^3)
4.9373 | well volume (L)
math check
24 | stage (in)
60.96 | stage (cm)
4937.2811 | well volume (cm^3)
4.9373 | well volume (L)
from transducer calibration
18 | stage (in)
45.72 | stage (cm)
3647.653 | drainage volume (cm^3)
22.2183 | drainage flux (mm)

## Counter Wt.
|  |  | SS rod dia (in) | SS rod pounds per inch
SS rod dia | 2 | (in) | 0.25 | 0.014
SS rod h | 28 | (in) | 1 | 0.229
SS rod wt | 0.917 | (pound/in) | 1.25 | 0.358
SS rod wt | 25.284 | (pound) | 1.5 | 0.516
|  |  | 1.75 | 0.702
|  |  | 2 | 0.917
|  |  | 2.25 | 1.161
|  |  | 2.5 | 1.433
|  |  | 2.75 | 1.734
|  |  | 3 | 2.064
|  |  | 4 | 3.669
|  |  | 5 | 5.733
|  |  | 6 | 8.256

## Transducer Calib
Transducer calibrations
Transducer was installed inside 4-inch ABS pipe allong with Turbo pump, cables and tygon tubbing
CS451
SN
|  | 20012329 | 20012339 | 21102340 | 20012326
| Volume (ml) | Stage (cm) | Stage (cm) | Stage (cm) | Stage (cm)
| 0 | 0 | 0 | 0 | 0
| 500 | 4.8 | 4.41 | 4.57 | 4.06
| 1000 | 12.05 | 11.63 | 11.87 | 11.45
| 1500 | 19.13 | 18.81 | 18.86 | 18.64
| 2000 | 25.47 | 25.15 | 25.26 | 24.91
| 2500 | 31.52 | 31.26 | 31.34 | 31.04
| 3000 | 37.59 | 37.19 | 37.49 | 37.08
| 3500 | 43.63 | 43.3 | 43.54 | 43.16
| 4000 | 49.76 | 49.43 | 49.73 | 49.39
average
| slope | 79.0753 | 79.2728 | 78.9597 | 79.1378 | 79.1114
| intercept | 32.3436 | 51.8277 | 46.5367 | 67.8937